Improvements for V0.3 of the Rastaban PCB

|  |  |  |
| --- | --- | --- |
| # | Improvement | Finished |
| 1 | Figure out which hardware pins to use for what components |  |
| 2 | TMC 2209 uses different ms1 ms2 configuration for microstepping than the tmc 2208! Keep this in mind. |  |
| 3 | Remove tmc2208 from design, focus on 2209 |  |
| 4 | Use UART on 2209 and remove step/dir enable interface. |  |
| 5 | Use appropriate resistor for UART control on 2209 (resistance should decrease with increase in drivers). |  |
| 6 | Connect diag pin of 2209 for stall (stuck motor) indication for rpi?? |  |
| 7 | Checkout the problems with diag pin on tmc2209 (see pdf in datasheets) |  |